## Definition phase
Scope of Analysis
Purpose statement and objectives | Identify disfunctions
Identify safety risks
Design proof, explore design robustness
Find opportunities for reliability improvement
Working conditions and scenarios under analysis | Scenario 1: magnetic field mapping of a permanent magnet for low field MRI
Scenario 2: sensor measurements in- and outside a high field MR scanner (B0 >= 1.5T)
Scenario 3: RF induced temperature/ time-varying electromagnetic field mapping for implant safety
Requirements | High fidelity measurements in quantity, location and time (including E-fields, H-fields and temperature)
Submillimeter resolution
Sufficient holding/ moving torque for a variety of probes
Reproducibility/ backlash-free performance
Robust configuration
Efficient use of financial resources/ low budget (2000€)
Large working volume, mapping volume
Open towards all directions to allow a flexible setting of the measurement probe (vertically/ horizontally)
Application with electromagnetic, temperature and ultrasound sensors
Support various of measurement probes
Easily modified and adapted to other functionalities
Low noise, smooth running
Transportable, movable

## Funcional & Failure Analysis
Functional and Failure Analysis
1st level | Function 1st level | Failures 1st level | 2st level | Function 2nd level | Failures 2nd level | 3rd level | Function 3rd level | Failures 3rd level | 4th level | Functions 4th level | Failures 4th level | 5th level | Functions 5th level | Failures 5th level | 6th level | Functions 6th level | Failures 6th level | 7th level | Functions 7th level | Failures 7th level
COSI Measure | working volume: (53x53x64)cm³ |  | mechanical subsystem | robust configuration | collapses | base frame | hold up to 55 kg (downwards) | bents | Aluminum profiles (40x80x800) | bottom x-axis and z-axis: low weight, strong, corrosion resistant, electrical conductivity, thermal conductivity durability, inductile, magnetic neutrality | dent, vibrates, crooked, creeping, deflection, fracture, overheated,
| precision: positioning error < 1mm | positioning error > 1mm |  | stable standing | skewed, wiggles, tipping over |  | stable standing, balance load | vibration, wiggles, imbalanced load, damages floor | Angle, 40x80 | Attachment Aluminum Frames X-axis to Z-axis bottom, 90 degree, apposite, non magnetic | deformed, magnetic acceleration,
| speed: 37 mm/s | slow movement |  | allows smooth and steady movement | stagnates, jerking movements |  | strong mechanical bond | collapses, crooked, skewed, loose connection, joint too large | screw M8x18 | Fixation Frame to angle (X-axis to Z-axis bottom) | loose, overtighten, driven in crooked, magnetic acceleration, tension, corrosion, embedment
| max load: 21 kg | system fails before max load reached |  | robust against temperature or magnetic field distribution (neutral) | influencing temperature distribution, influencing magnetic field distribution |  | withstand mechanical and electromagnetic stresses | magnetic | slot nut N8/ M8 | bolted joint Frame to angle | loose joint, fatigue failure
| applied force: 896 N (91 kg) |  |  | support submillimeter resolution | insufficient resolution |  | axially parallel | non axially parallel | Angle, 40 | Attachment Aluminum Frames Y axis to Z axis bottom | deformed, magnetic acceleration,
| high fidelity measurements in quantity, location and time (including E-fields, H-fields and temperature) | inaccurate measurements |  | transportable | not transportable |  |  | electrically conductive | screws M8x18, countersunk (or screws M8x12 as in BOM referenced?) | bolted joint angle to Frame (Y-axis to Z-axis bottom) | loose, overtighten, driven in crooked, magnetic acceleration, tension, corrosion, embedment
| Submillimeter resolution | insufficient resolution | [6 empty] | slot nut N8/ M8 | bolted joint angle to Frame Y axis to Z axis bottom, absorb mechanical forces | crack,
| reproducibility | low reproducibility, backlash | [6 empty] | Aluminum profiles (40x40x800) | Bottom y-axis: low weight, strong, corrosion resistant, durability, inductile, nonmagnetic, levelled surface for sufficient load bearing | creeping, vibrates, reformatted, fracture, overheated
| robust configuration | susceptible to outside influences, breaks under stress | [6 empty] | slot nut N8/ M8 | bolted joint angle to Frame: Y axis to Z axis and X axis to Z axis top, absorb mechanical forces
| Capability to position measurement along three dimension | no 3D-measurement possible | [6 empty] | slot nut N8/ M8 | bolted joint angle to Frame: X axis top, absorb mechanical forces
| safe usage | human injury due to mechanical force or electric shock/ burn | [6 empty] | Aluminum profiles (40x80x880) | X axis top: low weight, strong, corrosion resistant, electrical conductivity, thermal conductivity durability, inductile, nonmagnetic | creeping, vibrates, deformed, corroded, overheated
[9 empty] | Angle, 40x80 | Attachment Y axis to Z axis top
[9 empty] | screws, M8x18 | bolted joint angle to frame: X axis top | loose, overtighten, driven in crooked, magnetic acceleration, tension, corrosion, embedment
[9 empty] | Aluminum profiles (40x40x800) | Y axis top: low weight, strong, corrosion resistant, electrical conductivity, thermal conductivity durability, inductile, nonmagnetic | creeping, vibrates, deformed, fracture, overheated
[9 empty] | screws M8x18, countersunk (or screws M8x12 as in BOM referenced?) | bolted joint angle to frame: Y Axis to Z axis top (angle to aluminum frame) | loose, overtighten, driven in crooked, magnetic acceleration, tension
[9 empty] | slot nut N8/ M8 | bolted joint angle to frame Y axis to Z axis top
[9 empty] | Angle, 40 | Attachment Y axis to Z axis top
[6 empty] | 3-axis movement system | precise actuation of planned position | deviation, low reproducibility | X movement | moves smoothly and consistent in x- direction up to max load | vibration, clatters (conjunct movement) | linear axis (powertrain) | silent and efficient force transmission | squeaks, inefficient transmission | ball screw & spindle nut block | translate rotational movement to linear with high precision, little internal friction, drive heavy loads | high friction, uneven rotation, jammed, squeaks | ball screw | rotate evenly, precisely processed,
[7 empty] | efficient force transmission | loss of power |  | submillimeter linear resolution in x-direction | resolution too low |  | robust configuration | battered |  |  |  | spindle nut block | translate rotational movement to linear with high precision, little internal friction
[7 empty] | reproducible movement/ performance | functioning too slow/ slow transition of movements |  | high accuracy in X-direction | low accuracy, backlash |  | low wear | high wear |  |  |  | screw M5x20, countersunk | attach spindle nut block to ball screw, ensure force transmission under high load
[7 empty] | smooth running | clattering | [6 empty] | fixed bearing & motor holder | attach ball screw | offset, loose ball screw | Festlagerblock FLB20 / Easy-Mechatronics System 1620B, Art.Nr. 90043 | Allow only rotational movement, channel forces, | bearing failure
[7 empty] | submillimeter spatial resolution | spatial resolution too low | [7 empty] | align and hold motor | clearance | Lagerhalteplatte 62x62mm / Easy-Mechatronics System 1620B, Art.Nr. 90067
[7 empty] | align and adjust electrical subsystem | wrongly adjusted | [7 empty] | Allow only rotational movement, channel forces, | vertical or horizontal clearance | Abstandhülse d=5.5 ¦ D=10 ¦ L=25, Art.Nr. 70908
Motorplatte NEMA23 / Easy-Mechatronics System 1620A/1620B, ArtNr. 90029
DIN 912 Zylinderschraube mit Innensechskant, 8.8, verzinkt M5x22, Art,Nr. DIN912-5x22
DIN 912 Zylinderschraube mit Innensechskant, 8.8, verzinkt M5x35, Art,Nr. DIN912-5x35
Schrägkugellager zweireihig 3200/5200 2RS 10x30x14,3mm, Art.Nr. 32784
DIN 439 Sechskantmutter niedrige Form, verzinkt M10X1, Art.Nr. DIN439-M10
[15 empty] | floating bearing | Attachment ball screw | bearing failure | Loslagerblock LLB20 / Easy-Mechatronics System 1620B, Art.Nr. 90050
[16 empty] | only allow rotational movement | vertical or horizontal clearance | Rillenkugellager 6200 2RS 10x30x9mm, Art.Nr. 32593
[16 empty] | channel forces | friction | DIN 471 Sicherungsring für Wellen, Stahl, blank A10, Art.Nr. 71004
[15 empty] | motor coupler | Force transmission, resilience 5 NM, circular motion with angular accuracy | breaks burden too high
[12 empty] | attachment to base frame left | stable attachment | loose | slot nuts, M8 | bolted joint Plate 3 and plate 4 to frame (x/ y plane)
[13 empty] | parallel attachment to base frame | skewed | plate 3 | hold linear axis (x-axis) | breaks
[15 empty] | plate 4 | hold linear axis (x-axis) | breaks
[15 empty] | screws, M8x20, countersunk | Fixation Plate 3 and plate 4 to frame (x/ y plane) | loose, overtighten, driven in crooked, magnetic acceleration, tension, corrosion, embedment
[15 empty] | screws, M6x30 | bolted joint linear axis to plate 4 | loose, overtighten, driven in crooked, magnetic acceleration, tension, corrosion, embedment
[15 empty] | screw nuts M6 | bolted joint linear axis to plate 4 | loose, overtighten, driven in crooked, magnetic acceleration, tension, corrosion, embedment
[15 empty] | screwsM6x20 | bolted joint linear axis to plate 3 | loose, overtighten, driven in crooked, magnetic acceleration, tension, corrosion, embedment
[12 empty] | linear rail x direction | move smoothly in x - direction | friction, squeaks | slot nuts M5 | bolted joint supporting rail to top of base frame
[15 empty] | supporting rail, TBS20 | low friction guidance, high stiffness (rigidity), travel accuracy | deflection
[15 empty] | Screws, M5x20 | bolted joint supporting rail to top of base frame | loose, overtighten, driven in crooked, magnetic acceleration, tension, corrosion, embedment
[15 empty] | Linear Bearing, TBR20UU | move smoothly over supporting rail
[12 empty] | limit switch holder x-axis | hold limit switch | breaks
[13 empty] | adjust limit switch | wiggles, limit switch wrongly adjusted
[9 empty] | Y movement | move smoothly and consistent in Y- direction (backwards and forwards) | vibration, clatters (conjunct movement) | linear axis (powertrain) | silent and efficient force transmission | squeaks, inefficient transmission | ball screw & spindle nut block | move smoothly, translate rotational movement to linear with high precision, little internal friction, drive heavy loads | high friction, uneven rotation, jammed, squeaks | ball screw | rotate evenly, precisely processed,
[10 empty] | submillimeter linear resolution in Y-direction | resolution too low |  | robust configuration | battered |  |  |  | spindle nut block | translate rotational movement to linear with high precision, little internal friction
[10 empty] | high accuracy in Y-direction | low accuracy, backlash |  | low wear | high wear |  |  |  | screw M5x20, countersunk | bolted joint spindle nut block to ball screw, ensure force transmission under high load
[10 empty] | steady velocity | unsteady velocity |  |  |  | fixed bearing & motor holder | attach ball screw | offset, loose ball screw | Festlagerblock FLB20 / Easy-Mechatronics System 1620B, Art.Nr. 90043
[16 empty] | align and hold motor | clearance | Lagerhalteplatte 62x62mm / Easy-Mechatronics System 1620B, Art.Nr. 90067
[16 empty] | Allow only rotational movement | vertical or horizontal clearance | Abstandhülse d=5.5 ¦ D=10 ¦ L=25, Art.Nr. 70908
[16 empty] | channel forces |  | Motorplatte NEMA23 / Easy-Mechatronics System 1620A/1620B, ArtNr. 90029
DIN 912 Zylinderschraube mit Innensechskant, 8.8, verzinkt M5x22, Art,Nr. DIN912-5x22
DIN 912 Zylinderschraube mit Innensechskant, 8.8, verzinkt M5x35, Art,Nr. DIN912-5x35
Schrägkugellager zweireihig 3200/5200 2RS 10x30x14,3mm, Art.Nr. 32784
DIN 439 Sechskantmutter niedrige Form, verzinkt M10X1, Art.Nr. DIN439-M10
[15 empty] | floating bearing | Attachment ball screw | bearing failure | Loslagerblock LLB20 / Easy-Mechatronics System 1620B, Art.Nr. 90050
[16 empty] | only allow rotational movement | vertical or horizontal clearance | Rillenkugellager 6200 2RS 10x30x9mm, Art.Nr. 32593
[16 empty] | channel forces | friction | DIN 471 Sicherungsring für Wellen, Stahl, blank A10, Art.Nr. 71004
[15 empty] | motor coupler | Force transmission, resilience 5 NM, circular motion with angular accuracy | breaks (burden too high)
[12 empty] | fixation of linear axis (y) | stabilize linear axis (y) |  | plate 1 | translate x-movement to axis/ stabilize linear y axis
[13 empty] | align axis and rail y-direction | non parallel alignment | screws, M6x20 countersunk | bolted joint: plate 1 to four linear bearings (plate 1 to linear bearings on x axis) | loose, overtighten, driven in crooked, magnetic acceleration, tension, corrosion, embedment
[15 empty] | screws, M6x20 | bolted joint; plate 1 to linear axis | loose, overtighten, driven in crooked, magnetic acceleration, tension, corrosion, embedment
[12 empty] | linear rail y direction | move smoothly in y- direction | squeaks | supporting rail | low friction guidance, high stiffness (rigidity), travel accuracy, hold z axis | friction, deformation under load
[13 empty] | hold up to 55,1 kg | bents | linear bearing | move smoothly over supporting rail, pass forces of y-movement to plate 5,6 and 7 | crack
[15 empty] | screws M5x30 | bolted joint linear bearing to plate 1 | loose, overtighten, driven in crooked, magnetic acceleration, tension, corrosion, embedment
[15 empty] | Hexagon nut, M5 | fix supporting rails to plate 1 | loose, skewed attachment to screw,
[12 empty] | attachment to x-axis | force transmission of x-axis to linear rail (x-direction) | inefficient transmission, no transmission | plate 2 | Attachment plate 1 to ball screw block from the x axis, force transmission | crack, deformation
[15 empty] | screws M5x30 | bolted joint plate 2 and plate 1 | loose, overtighten, driven in crooked, magnetic acceleration, tension, corrosion, embedment
[15 empty] | screws M5x20, countersunk | bolted joint ball screw block and plate 2 | loose, overtighten, driven in crooked, magnetic acceleration, tension, corrosion, embedment
[15 empty] | Distance plate 1 mm | tight fit | clearance
[12 empty] | limit switch holder y-axis | hold limit switch | breaks
[13 empty] | adjust limit switch | wiggles, limit switch wrongly adjusted
[9 empty] | Z movement | move smoothly and consistent in Z- direction (backwards and forwards) | vibration, clatters (conjunct movement) | linear axis | silent and efficient force transmission | squeaks, inefficient transmission | Easy mechatronic system | silent and efficient force transmission | squeaks, inefficient transmission | fixed bearing & motor holder
[10 empty] | submillimeter linear resolution in Z-direction | resolution too low |  | robust configuration | battered |  | robust configuration | battered | floating bearing
[10 empty] | high accuracy in Z-direction | low accuracy, backlash |  | low wear | high wear |  | low wear | high wear | linear bearings SCJ20UU
[10 empty] | steady velocity | unsteady velocity |  | hold and stabilize probe | probe wiggles, probe wrongly adjusted |  |  |  | präzisionshaft 20mm h6 840mm
shaft holder SH20
ball screw with spindle nut block SFU1605-DM 852mm
shaft coupling 8/8mm D25L30
[15 empty] | plate 8 | align Easy- Mechatronics system, stabilize | skewed alignment
[15 empty] | screws, M6x20 | bolted joint Easy- Mechatronics system to plate 8 and 9 | loose, overtighten, driven in crooked, magnetic acceleration, tension, corrosion, embedment
[15 empty] | plate 9 | hold probe | crack, deformed, bent, overheat
[12 empty] | plate assembly | force transmission of y-axis to z-axis | force loss | plate 5 | attach z-axis to upper linear bearings, robust configuration strong,, thermal conductivity durability, inductile, nonmagnetic | crack, deformed, bent, overheat
[15 empty] | plate 6 | attach z-axis to plate 5 and 7, robust configuration, thermal conductivity, durability, nonmagnetic | crack, deformed, bent, overheat
[15 empty] | plate 7 | attach z-axis to lower linear bearings, robust configuration, thermal conductivity, inductile, nonmagnetic | crack, deformed, bent, overheat
[15 empty] | screws, M5x30 (4x) | Connect plate 5 with plate 6 and plate 6 with plate 7
[15 empty] | screws, M6x20, countersunk | bolted joint: plates to linear bearing on supporting rail of easy mechatronics system
[15 empty] | screws, M5x20, countersunk | bolted joint plate 6 to ball screw block on z-axis
[12 empty] | attachment to y-axis | Attach easy mechatronics system to supporting rails and motor holder | skewed attachment, loose attachment | screws, M6x20, countersunk | bolted joint: plates to linear bearing on y-axis | loose, overtighten, driven in crooked, magnetic acceleration, tension, corrosion, embedment
[15 empty] | screws, M5x20, countersunk | bolted joint plate 6 to ball screw block on y-axis | loose, overtighten, driven in crooked, magnetic acceleration, tension, corrosion, embedment
[15 empty] | distance plate, 1 mm | reduce clearance between linear bearings and/ or plate 5 and plate 7
[12 empty] | limit switch holder z-axis | hold limit switch | breaks
[13 empty] | adjust limit switch | wiggles, limit switch wrongly adjusted
[6 empty] | moving table | roll smoothly | does not roll, wiggles,
[7 empty] | hold total system | collapses
[7 empty] | provide stable underground | skewed, rolls unintentionally

## FMEA
FMEA Form Sheet
|  | Project: |  | Design FMEA x | [6 empty] | Process FMEA
|  | COSI Measure |  | Name/ Community: |  | Process/ Product name: |  |  |  |  | Contributor: | Email Address | Modified date | Commit Message
|  |  |  | Open Source Imaging |  | COSI Measure mechanical subsystem |  |  |  |  | Researcher A |  | 2019-01-13 | Initial Commit
[11 empty] | Participant B |  | 2019-01-31 | Limit switch and mobile wagon added to structure analysis
[11 empty] | Participant C |  | 2019-02-15 | Risk Assessment
[11 empty] | Participant A |  | 2019-04-12 | Scenarios adjusted, risk assessment changed
[11 empty] | Participant B |  | 2019-04-17 | Risk Assessment
[6 empty] | Status quo | [6 empty] | Improved Status
| ID | Position/ process step | Potential failure mode | Failure effect | Failure cause | Preventive Actions | O* | S* | D* | RPN* | Optimization Actions | Responsible | Implemented measures | O* | S* | D* | RPN*
| 34. | base frame | base frame carries electric current | serious human injury | electrically conductive components (aluminum) |  | 8 | 10 | 6 | 480 | grounded base frame
| 129. | mechanical subsystem | electro-magnetic induction | high measurements influences on high frequency electromagnetic field | electrically conductive components |  | 8 | 8 | 6 | 384 | application specific failure. longer probe holder, shielding
| 116. | limit switch holder (y-axis) | breaks | limit switch inoperative | fragile 3D printed parts |  | 7 | 7 | 6 | 294 | stronger material e.g. SLA 3D prints instead of filament 3D prints, additional emergency stop sensors (e.g. pressure)
| 123. | limit switch holder (z-axis) | breaks | limit switch inoperative | fragile 3D printed parts |  | 7 | 7 | 6 | 294 | stronger material e.g. SLA 3D prints instead of filament 3D prints, additional emergency stop sensors (e.g. pressure)
| 97. | limit switch holder (x-axis) | breaks | limit switch inoperative | fragile 3D printed parts |  | 7 | 7 | 6 | 294 | stronger material e.g. SLA 3D prints instead of filament 3D prints, additional emergency stop sensors (e.g. pressure)
| 35. | 3-axis movement system | probe bumps into object | axis is damaged | no force feedback/ reactivity |  | 6 | 8 | 5 | 240 | monitor current and give feedback when something was hit.
| 99. | linear axis (powertrain) x-axis | plate 2 keeps pushing against fixed bearing and motor holder | motor is damaged | limit switch inoperative |  | 5 | 8 | 5 | 200
| 39. | 3-axis movement system | deviation | inaccurate measurements | backlash in z-direction |  | 5 | 7 | 5 | 175
| 33. | base frame | sharp underground | damages floor | sharp aluminum profiles |  | 7 | 7 | 3 | 147
| 32. | base frame | high vibrations | slow performance (time to measure) | high magnetic field |  | 6 | 6 | 4 | 144
| 37. | 3-axis movement system | deviation | inaccurate measurements | backlash in X direction |  | 4 | 7 | 5 | 140
| 38. | 3-axis movement system | deviation | inaccurate measurements | backlash in Y direction |  | 4 | 7 | 5 | 140
| 55. | X-movement | backlash | deviation in x- direction | ball screw resolution too low |  | 4 | 7 | 5 | 140
| 65. | Y-movement | backlash | deviation in y- direction | ball screw resolution too low |  | 4 | 7 | 5 | 140
| 75. | Z-movement | backlash | deviation in z- direction | ball screw resolution too low |  | 4 | 7 | 5 | 140
| 56. | X-movement | linear resolution too low | insufficient measurement resolution | ball screw resolution too low |  | 4 | 7 | 5 | 140
| 66. | Y-movement | linear resolution too low | insufficient measurement resolution | ball screw resolution too low |  | 4 | 7 | 5 | 140
| 76. | Z-movement | linear resolution too low | insufficient measurement resolution | ball screw resolution too low |  | 4 | 7 | 5 | 140
| 22. | base frame | non axially parallel | skewed mechanical subsystem | wrong cut length of aluminum profiles |  | 6 | 7 | 3 | 126 | measure parts when supply arrives
| 134. | base frame | enclosure for electronics comes loose | damages electronics | attachment is weak |  | 2 | 5 | 3 | 30 | redesign the attachment
| 135. | base frame | enclosure for electronics comes loose | damages wiring | attachment is weak |  | 2 | 5 | 3 | 30 | redesign the attachment
| 15. | mechanical subsystem | static magnetic field | inaccurate measurements of magnetic field | ferromagnetic components |  | 4 | 5 | 6 | 120 | coverage, background scan
| 105. | linear axis (powertrain) y-axis | no force transmission | linear y-axis inoperative | fixed bearing failure |  | 4 | 6 | 5 | 120
| 52. | X-movement | clatters | loss of force | inefficient transmission of linear axis |  | 4 | 6 | 5 | 120
| 53. | X-movement | clatters | loss of force | high wear of linear axis |  | 4 | 6 | 5 | 120
| 54. | X-movement | clatters | loss of force | skewed attachment of linear axis |  | 4 | 6 | 5 | 120
| 87. | linear axis (powertrain) x-axis | no force transmission | linear x-axis inoperative | fixed bearing failure |  | 4 | 6 | 5 | 120
| 36. | 3-axis movement system | uncoordinated movements between three axis | irregular movement | limit switches wrongly adjusted |  | 4 | 7 | 4 | 112
| 43. | 3-axis movement system | insufficient measurement resolution | customer requirement not achieved | linear resolution in X- direction too low |  | 4 | 7 | 4 | 112
| 44. | 3-axis movement system | insufficient measurement resolution | customer requirement not achieved | linear resolution in Y- direction too low |  | 4 | 7 | 4 | 112
| 45. | 3-axis movement system | insufficient measurement resolution | customer requirement not achieved | linear resolution in Z- direction too low |  | 4 | 7 | 4 | 112
| 30. | base frame | vibrates | slow performance (time to measure) | bad probe attachment (e.g. long rod) |  | 6 | 6 | 3 | 108 | not low reproducibility but time to measure
| 1. | mechanical subsystem | tipping over | minor human injury | imbalanced load |  | 4 | 9 | 3 | 108 | document the importance of a stable standing and the potential danger
| 4. | mechanical subsystem | tipping over | loss of system | imbalanced load |  | 3 | 9 | 4 | 108 | Document maximum load and the potential danger that can occur if the system collapses
| 24. | base frame | non axially parallel | minor skewed mechanical subsystem | overtight screws |  | 2 | 5 | 6 | 60
| 101. | linear axis (powertrain) y-axis | high wear | y-movement clatters | high friction of spindle nut block |  | 3 | 6 | 6 | 108
| 102. | linear axis (powertrain) y-axis | high wear | y-movement clatters | high friction in floating bearing |  | 3 | 6 | 6 | 108
| 103. | linear axis (powertrain) y-axis | high wear | y-movement clatters | high friction in fixed bearing |  | 3 | 6 | 6 | 108
| 83. | linear axis (powertrain) x-axis | high wear | x-movement clatters | high friction of spindle nut block |  | 3 | 6 | 6 | 108
| 84. | linear axis (powertrain) x-axis | high wear | x-movement clatters | high friction in floating bearing |  | 3 | 6 | 6 | 108
D | 85. | linear axis (powertrain) x-axis | high wear | x-movement clatters | high friction in fixed bearing |  | 3 | 6 | 6 | 108
| 89. | linear axis (powertrain) x-axis | squeaks | noisy movement | offset of ball screw to fixed bearing |  | 3 | 6 | 6 | 108
| 104. | linear axis (powertrain) y-axis | limited force transmission | movement in y-direction limited | floating bearing failure |  | 4 | 5 | 5 | 100
| 86. | linear axis (powertrain) x-axis | limited force transmission | movement in x-direction limited | floating bearing failure |  | 4 | 5 | 5 | 100
| 72. | Z-movement | vibrates | slow performance (time to measure) | load too high | Max load is stated (tested/ calculated?) | 4 | 6 | 4 | 96 | potentially test/ calculate max load
| 100. | linear axis (powertrain) y-axis | high wear | y-movement clatters | insufficient configuration for load |  | 4 | 6 | 4 | 96
| 82. | linear axis (powertrain) x-axis | high wear | x-movement clatters | insufficient configuration for load |  | 4 | 6 | 4 | 96
| 49. | X-movement | vibrates | slow performance (time to measure) | load too high | Max load is documented | 4 | 6 | 4 | 96 | potentially test/ calculate max load
| 62. | Y-movement | vibrates | slow performance (time to measure) | load too high | Max load is documented | 4 | 6 | 4 | 96 | potentially test/ calculate max load
| 132. | linear rail x-direction | squeaks | Customer satisfaction reduced | load too high |  | 4 | 6 | 4 | 96
| 110. | linear rail y-direction | squeaks | Customer satisfaction reduced | load too high | Max load is documented | 4 | 6 | 4 | 96 | potentially test/ calculate max load
| 107. | linear axis (powertrain) y-axis | squeaks | noisy movement | offset of ball screw to fixed bearing |  | 3 | 6 | 5 | 90
| 108. | linear axis (powertrain) y-axis | inefficient transmission | y-movement clatters | clearance of motor and motor holder |  | 3 | 6 | 5 | 90
| 115. | Attachment to x-axis | force loss between x-axis and linear rail (x-direction) | x-movement clatters | plate 2 worn out |  | 3 | 6 | 5 | 90
| 131. | linear rail x-direction | squeaks | Customer satisfaction reduced | linear bearing too stiff |  | 3 | 6 | 5 | 90
| 7. | mechanical subsystem | wiggles | positioning error | base frame: imbalanced load |  | 4 | 7 | 3 | 84
| 8. | mechanical subsystem | wiggles | low reproducibility | base frame: imbalanced load |  | 4 | 7 | 3 | 84
| 58. | X-movement | low repeatability x-axis | low reproducibility | motor error too high |  | 2 | 7 | 6 | 84
| 59. | X-movement | low accuracy x-axis | reduced measurement accuracy | bad motor linearity |  | 2 | 7 | 6 | 84
| 60. | X-movement | low accuracy x-axis | reduced measurement accuracy | motor error too high |  | 2 | 7 | 6 | 84
| 68. | Y-movement | low repeatability y-axis | low reproducibility | motor error too high |  | 2 | 7 | 6 | 84
| 70. | Y-movement | low accuracy y-axis | reduced measurement accuracy | bad motor linearity |  | 2 | 7 | 6 | 84
| 71. | Y-movement | low accuracy y-axis | reduced measurement accuracy | motor error too high |  | 2 | 7 | 6 | 84
| 78. | Z-movement | low accuracy z-axis | reduced measurement accuracy | motor error too high |  | 2 | 7 | 6 | 84
| 80. | Z-movement | low accuracy z-axis | reduced measurement accuracy | bad motor linearity |  | 2 | 7 | 6 | 84
| 81. | Z-movement | low repeatability z-axis | low reproducibility | motor error too high |  | 2 | 7 | 6 | 84
| 2. | mechanical subsystem | tipping over | minor human injury | load severely too high | Max load is documented | 3 | 9 | 3 | 81 | Document maximum load and the potential danger that can occur if the system collapses
| 5. | mechanical subsystem | tipping over | loss of system | load severely too high | Max load is documented | 3 | 9 | 3 | 81 | Document maximum load and the potential danger that can occur if the system collapses
| 19. | base frame | crooked | system is not functioning | load severely too high | Max load is documented | 3 | 9 | 3 | 81 | Document maximum load and the potential danger that can occur if the system collapses
| 3. | mechanical subsystem | tipping over | minor human injury | base frame collapses | more angle, screws and joints were added to the design | 3 | 9 | 3 | 81 | Document maximum load and the potential danger that can occur if the system collapses
| 113. | Attachment to x-axis | no force transmission between x-axis and linear rail (x-direction) | no movement in x-direction | plate 2 breaks |  | 2 | 8 | 5 | 80
| 29. | base frame | magnetic | inaccurate measurements/ deviation | wrong material supply | link for correct material and supplier available | 2 | 8 | 5 | 80
| 109. | fixation of linear y-axis | friction | transmission loss y-direction | non-parallel alignment of y-axis and y-rail |  | 4 | 6 | 3 | 72
| 27. | base frame | joint too large | loose connection | uneven tightening torque |  | 4 | 6 | 3 | 72
| 126. | moving table | collapses | system damaged | load severely too high | Max load is documented | 3 | 8 | 3 | 72
| 12. | mechanical subsystem | jerking movements | positioning error | base frame non axially parallel assembled |  | 3 | 8 | 3 | 72
| 120. | attachment to y-axis | loose attachment | backlash | clearance between plate 7 and bearing | distance plate 1mm | 3 | 8 | 3 | 72
| 13. | mechanical subsystem | jerking movements | low reproducibility | base frame non axially parallel assembled |  | 3 | 8 | 3 | 72
| 18. | linear axis z-axis | probe wiggles | inaccurate measurements | loose attachment plate 9 to easy-mechatronics system | more angle, screws and joints were added to the design | 3 | 8 | 3 | 72
| 63. | base frame | collapses | minor human injury | loose screws |  | 3 | 8 | 3 | 72
| 73. | base frame | wiggles | unstable standing of mechanical system | loose screws |  | 3 | 8 | 3 | 72
| 94. | attachment to base-frame left x-axis | loose | x-movement backlash | torque too low |  | 3 | 8 | 3 | 72
| 114. | Attachment to x-axis | force loss between x-axis and linear rail (x-direction) | x-movement vibration | clearance | distance plate 1mm | 3 | 6 | 4 | 72
| 96. | linear rail x-direction | friction | x-movement clatters | rails are not parallelly assembled |  | 3 | 6 | 4 | 72
| 23. | base frame | non axially parallel | minor skewed mechanical subsystem | wrong assembly of angles and profiles |  | 3 | 6 | 4 | 72
| 88. | linear axis (powertrain) x-axis | inefficient transmission | movement in x-direction slow | uneven rotation of ball screw and spindle nut block |  | 3 | 6 | 4 | 72
| 57. | X-movement | linear resolution too low | insufficient measurement resolution | motor resolution too low |  | 2 | 7 | 5 | 70
| 67. | Y-movement | linear resolution too low | insufficient measurement resolution | motor resolution too low |  | 2 | 7 | 5 | 70
| 77. | Z-movement | linear resolution too low | insufficient measurement resolution | motor resolution too low |  | 2 | 7 | 5 | 70
| 106. | linear axis (powertrain) y-axis | inefficient transmission | movement in y-direction slow | uneven rotation of ball screw and spindle nut block |  | 3 | 5 | 4 | 60
| 90. | linear axis (powertrain) x-axis | inefficient transmission | x-movement clatters | clearance of motor and motor holder |  | 2 | 6 | 5 | 60
| 130. | linear rail x-direction | bents under load | x-movement clatters | load severely too high |  | 3 | 6 | 3 | 54
| 111. | linear rail y-direction | bents under load | y-movement clatters | load severely too high | Max load is documented | 3 | 6 | 3 | 54 | potentially test/ calculate max load
| 6. | mechanical subsystem | tipping over | loss of system | base frame collapses |  | 3 | 9 | 2 | 54 | Document maximum load and the potential danger that can occur if the system collapses
| 118. | limit switch holder (x-axis) | wiggles | limit switch function limited | loosely fixed to axis |  | 3 | 6 | 3 | 54
| 119. | limit switch holder (y-axis) | wiggles | limit switch function limited | loosely fixed to axis |  | 3 | 6 | 3 | 54
| 21. | base frame | skewed | jerking movements | bunted aluminum profiles |  | 3 | 6 | 3 | 54
| 91. | attachment to base-frame left x-axis | loose | x-movement clatters | torque too low |  | 3 | 6 | 3 | 54
| 95. | attachment to base-frame left x-axis | skewed | x-movement clatters | wrong assembly of plates |  | 3 | 6 | 3 | 54
| 117. | base frame | non axially parallel | minor skewed mechanical subsystem | loose screws |  | 3 | 6 | 3 | 54
| 124. | limit switch holder (z-axis) | wiggles | limit switch function limited | loosely fixed to axis |  | 3 | 6 | 3 | 54
| 112. | linear rail y-direction | squeaks | Customer satisfaction reduced | linear bearing too stiff |  | 2 | 5 | 5 | 50
| 121. | attachment to y-axis | skewed attachment | inaccurate positioning | plate assembly offset |  | 3 | 8 | 2 | 48
| 9. | mechanical subsystem | wiggles | positioning error | base frame skewed |  | 2 | 8 | 3 | 48
| 46. | 3-axis movement system | low system accuracy | customer requirement not achieved | low accuracy of X - movement |  | 2 | 6 | 4 | 48
| 47. | 3-axis movement system | low system accuracy | customer requirement not achieved | low accuracy of Y - movement |  | 2 | 6 | 4 | 48
| 48. | 3-axis movement system | low system accuracy | customer requirement not achieved | low accuracy of Z - movement |  | 2 | 6 | 4 | 48
| 122. | plate assembly | force loss between y-axis and z-axis | max load cannot be carried | inefficient configuration | configuration was tested | 3 | 5 | 3 | 45
| 125. | moving table | wiggles | no stable standing of mechanical subsystem | wrong assembly |  | 3 | 7 | 2 | 42
| 128. | moving table | skewed | no stable standing of mechanical subsystem | wrong assembly |  | 3 | 7 | 2 | 42
| 127. | moving table | skewed | jerking movements in mechanical subsystem | wrong assembly |  | 3 | 6 | 2 | 36
| 40. | 3-axis movement system | clattering | inaccurate measurements | x movement clatters (conjunct movement) |  | 5 | 7 | 1 | 35
| 41. | 3-axis movement system | clattering | inaccurate measurements | y movement clatters (conjunct movement) |  | 5 | 7 | 1 | 35
| 42. | 3-axis movement system | clattering | inaccurate measurements | z movement clatters (conjunct movement) |  | 5 | 7 | 1 | 35
| 16. | mechanical subsystem | jerking movements | inaccurate measurements | movement system clatters |  | 5 | 7 | 1 | 35
| 92. | attachment to base-frame left x-axis | plate 3 breaks | no working attachment x-axis | Aluminum too soft/ design wrong |  | 1 | 7 | 5 | 35
| 93. | attachment to base-frame left x-axis | plate 4 breaks | no working attachment x-axis | Aluminum too soft/ design wrong |  | 1 | 7 | 5 | 35
| 10. | mechanical subsystem | no stable standing | positioning error | moving table skewed |  | 2 | 8 | 2 | 32
| 17. | mechanical subsystem | jerking movements | minor measurements impact, | moving table skewed |  | 2 | 7 | 2 | 28
| 20. | linear axis z-axis | probe wiggles | inaccurate measurements | loose attachment probe to probe holder | more angle, screws and joints were added to the design | 3 | 7 | 1 | 21
| 25. | X-movement | vibrates | low reproducibility | loose attachment to base frame | more angle, screws and joints were added to the design | 3 | 7 | 1 | 21
| 26. | Y-movement | vibrates | low reproducibility | loose attachment to base frame | more angle, screws and joints were added to the design | 3 | 7 | 1 | 21
| 50. | Z-movement | vibrates | low reproducibility | loose attachment to base frame |  | 3 | 7 | 1 | 21
| 11. | mechanical subsystem | no stable standing | positioning error | moving table wiggles |  | 2 | 8 | 1 | 16
| 133. | mechanical subsystem | influences measurement (temperature) | inaccurate temperature measurements | thermal conduction or temp. gradients |  | 1 | 2 | 5 | 10
O* ... Occurrence
probability that a failure cause will occur
very unlikely = 1
low = 2 - 3
moderate = 4 - 6 high = 7 - 8
very high = 9 - 10
S* ... Severity
Consequences of the failure effect for the customer
insignificant failure = 1
relatively insignificant failure = 2 - 3
relatively severe failure = 4 - 6
severe failure = 7 - 8
extremely severe failure = 9 - 10
D* ... Detection
Probability of detecting a failure cause prior to delivery
very probable = 1 probable = 2 - 3
moderate = 4 - 6
low = 7 - 8
very low = 9 - 10
RPN* ... Risk Priority Number
high <= 1000
moderate <= 250
low <= 125
no risk = 1